## Slide 1: Development of a Dual-sensor Wearable Phonocardiogram and Application of Signal Processing Techniques for Heart Sound Denoising and Segmentation
Yui Harayama
MEng Biomedical Engineering - Imperial College London
Supervisors: Dr Firat Güder and Dr Yasin Çotur

## Slide 2
Cardiovascular diseases are the leading cause of death globally [1]
Comparison of existing diagnostic methods:
Diagnosis Method | Characteristics | 
Electrocardiography (ECG) | Unable to detect murmurs and cumbersome process [2] | 
Auscultation | Limited by the human audition and depends on listening abilities of the physician [3] | 
Phonocardiography (PCG) | Easy process and provides an objective assessment of complex heart sounds [4] | 
1/19
Motivation of the Project
Notes: The most used existing diagnosis methods are ECG
Therefore, pcg which overcomes these challenges was identified as a promising alternative

## Slide 3
PCG: plot of the recording of the sound and murmurs produced by the mechanical activity of the heart. Recordings are achieved using microphone sensors.
Systole
Diastole
Fundamental heart sounds (FHS): S1 and S2 produced by closure of cardiac valves
Selected Method: Phonocardiography
Note:
PCG is sensitive to noise  requires specifically-designed wearable for signal acquisition and signal processing for the denoising of noisy heart sounds.
Lack of research focused on multi-sensor-based denoising algorithms  address this gap
2/19
Notes: More specifically ‘dual-sensor’

## Slide 4
Aim and Objectives
Aim:
To develop a wearable dual-sensor system to perform PCG in uncontrolled environments.
Objectives:
1. Design a wearable phonocardiogram building upon research group’s work
2. Tuning of the Denoising Algorithms’ Parameters using Simulated Data.
  Compare the dual-sensor-based Adaptive Noise Cancellation (ANC) and Dual Spectral Subtraction (DSS) with the widely used single-sensor-based Discrete Wavelet Transform.
ANC: Utilised by a few research, but denoised simulated heart sounds contaminated with White Gaussian Noise (WGN).
DSS: Only utilised by one research. Used real-world heart sounds but reported to achieve ‘moderate success’ without further specification [5].
3/19

## Slide 5
Aim:
To develop a wearable dual-sensor system to perform PCG in uncontrolled environments.
Objectives:
1. Design a wearable phonocardiogram building upon research group’s work
3. Denoising of Real-world Heart Sounds Using the Optimised Algorithms.
  Heart sounds are acquired using the proposed dual-sensor system.
4. Segmentation of Denoised Heart Sounds.
Aim and Objectives
2. Tuning of the Denoising Algorithms’ Parameters using Simulated Data.
  Compare the dual-sensor-based Adaptive Noise Cancellation (ANC) and Dual Spectral Subtraction (DSS) with the widely used single-sensor-based Discrete Wavelet Transform.
4/19

## Slide 6: Methods: Overview
Simulated heart sounds contaminated with noise (WGN)
Tuning of the denoising algorithms’ parameters
Optimal parameters of each algorithm
Objective 2
Acquire heart sounds using dual-sensor wearable
Objective 3
Denoising Algorithms
P
R
P
R
Methods: Overview
5/19
Objective 4
Denoised heart sounds
Segmentation Algorithm
S1
S2
Sys
Dias
Notes: the simulated heart sounds contaminated with noise are used to tune the algorithms’ parameters. Optimal parameters are then used to denoise real world heart sounds acquired using our proposed dual sensor wearable. Two sensors are used; the primary sensors collects the noise contaminated heart sound while the reference sensor collects noise only. Denoised heart sounds are then segmented to obtain s1 s2, systole and diastole.

## Slide 7
Methods: Adaptive Noise Cancellation
Primary signal:
Reference signal:
Estimated primary noise:
Error signal:
Heart sound:
Primary noise:
Inputs: and
Filter estimates
Output:
Update the filter to approach
ANC in time domain
Update
6/19

## Slide 8
Methods: Adaptive Noise Cancellation
Primary signal:
Reference signal:
Estimated primary noise:
Error signal:
Heart sound:
Primary noise:
ANC in time domain
Update
7/19
Notes: Such that estimated primary noise is close to the primary noise

## Slide 9
Methods: Adaptive Noise Cancellation
Primary signal:
Reference signal:
Estimated primary noise:
Heart sound:
Primary noise:
ANC in time domain
Denoised signal:
Update
8/19

## Slide 10
Methods: Dual Spectral Subtraction
Primary signal:
Reference signal:
Denoised signal:
Attenuation factor: controls the attenuation strength
DSS in frequency domain
S1
S2
9/19
Notes: product of the noisy signal spectrum and the frequency response of a spectral subtraction filter

## Slide 11
Methods: Dual Spectral Subtraction
Primary signal:
Reference signal:
Denoised signal:
DSS in frequency domain
S1
S2
This is the parameter that we are tuning in DSS by denoising simulated heart sounds contaminated with WGN
Values: [0.5, 0.6, 0.8, 1, 1.2, 2, 4, 6, 8, 10]
Attenuation factor: controls the attenuation strength
10/19
Notes: product of the noisy signal spectrum and the frequency response of a spectral subtraction filter

## Slide 12
Inverse Fourier Transform
DSS in frequency domain
Methods: Dual Spectral Subtraction
Primary signal:
Reference signal:
Denoised signal:
11/19
S1
S2

## Slide 13
Methods: Segmentation
Extract the signal’s envelope
2. Smooth the envelope
3. Identify the FHS
4. Identify the boundaries of the FHS
12/19

## Slide 14
Results: Parameter Tuning (DSS)
Objective 2
SI-SDR: power of over power of distortion
MSPE: average of the squares of the error between and
Original + WGN
Original
Denoised
Observations: There is a trade-off between effective removal of noise and attenuation of cardiac sounds
Lower = better
Higher = better
13/19
Notes: Here are the plots of the simulation of the noisy heart sounds in blue, the original clean heart sound in red and the denoised heart sound in yellow. The performance measurement metrics, the MSPE and SI-SDR were used to identify the optimal lambda. The lower the MSPE and the higher the SISDR the better the performance, and the highest performance was measured with lambda = 1. As you can see, small lambda results in a noisy denoised waveform but large lambda results in the attenuation of the S2. Therefore, there is a trade off between effective…

## Slide 15
Comparison of denoising algorithms’ average performance after tuning:
SI-SDR gain: how much the denoising has improved the SI-SDR
ANC
Algorithms | Superiority of DSS Over the Algorithm (in %)
LMS | 52.94
NLMS | 43.12
DWT | 40.54
Average | 45.53
Observations:
DSS outperforms the other algorithms
Requires and to be correlated
Results: Denoising (post tuning)
Objective 3
Least Mean Square
Normalised Least Mean Square
14/19
n = 5
n = 6
n = 7
n = 7
n = 4
n = 3
n = 3
Notes: Heart sounds from a total of 11 participants were acquired in environments contaminated with 7 noise sources. The si-sdr gain, which measures how much the denoising improved the signal quality was used to measure the performance of all algorithms. LMS and NLMS are two implementations of the ANC. On average, DSS outperforms other algorithms as indicated by the green arrows. Specifically, it outperforms on average the other algorithms by 45%. However, as you can see at the red arrow, DSS is outperformed with the body movement noise. This can be explained by some cases where the noise recorded by the primary and reference sensors were not correlated.

## Slide 16
Comparison of waveform of denoised heart sounds:
Observations: Heart sounds denoised with DSS are clearer than others
Objective 3
Results: Denoising (post tuning)
15/19
Notes: In this example, recordings were achieved in a busy café. It demonstrates how the waveform denoised using DSS is much clearer than others. Despite the heart sounds being not identifiable in the primary signal, s1 and s2 are clearly visible in the DSS waveform.

## Slide 17
Peak Correction: to rectify the delay between ground truth peaks and detected peaks
Results: Segmentation
Observations: Time misalignment is corrected
Start of FHS
End of FHS
Objective 4
16/19
Notes: The fhs were correctly identified, as indicated by the red circles, as well as the start and end boundaries indicated by the blue and green circles. However, a slight time misalignment of the fhs was observed, which was rectified after applying peak correction. The circles align with the maxima of the fhs.

## Slide 18
x
Δ
Peak Conditioning: to identify any missed peaks and remove false peaks
Results: Segmentation
Observations: Candidate peak is identified and extra peak is suggested to be removed
Objective 4
17/19
Notes: Peak conditioning to identify any missed peaks and remove false peaks was then applied during a second pass of the segmentation. A low amplitude s2 which was missed during the first pass is now identified as a candidate peak, while an extra peak due to an artefact is now suggested to be removed.

## Slide 19
Results: Segmentation
Observations:
i) In general, start and end boundaries are better identified than the peaks
ii) Peak-conditioning improves the performance
Objective 4
Metric (%) | Peaks (pre-conditioning) | Peaks (post-conditioning) | Start | End
Sensitivity | 52.94 | 82.70 | 91.41 | 91.41
Positive Predictivity | 43.12 | 98.15 | 83.49 | 98.81
Accuracy | 77.54 | 82.70 | 90.79 | 90.35
Detection Error Rate | 40.54 | 22.44 | 10.61 | 11.21
Average | 58.27 | 85.28 | 88.77 | 92.34
Higher = better
Lower = better
18/19
Notes: Finally, the segmentation performance was measured using these four metrics. In general, the start and end boundaries are better identified that the peaks as indicated by the higher average and peak conditioning improves the performance by 27%.

## Slide 20
Objective 2: Tuning of the Parameters
Attenuation factor must be carefully selected to effectively remove noise without
attenuating cardiac sounds.
Especially important with small amplitude heart sounds or murmurs.
Future work: tune with noise other than WGN.
Conclusion
Objective 3: Denoising of Real-world Heart Sounds
Addressed the gap in research of multi-sensor-based denoising algorithms.
DSS successfully outperformed other algorithms by an average of 45.53%.
Future work: use a larger and more varied subject base.
Objective 4: Segmentation of Denoised Heart Sounds
Segmentation algorithm successfully identified the FHS and boundaries.
Peak conditioning improved the performance by 27.01%.
However, performance is below the literature  use hidden Markov model.
19/19
Notes: This is especially important with small amplitude heart sounds or murmurs, which are indicators of cardiac pathologies as heavy attenuation of murmurs might lead to misdiagnosis.
To further refine the selection of λ, we can tune it based on denoising HS contaminated with everyday noise sources instead of solely relying on WGN.
Use alternative such as the HMM which is considered to be a state of the art segmentation algorithm

## Slide 21
[1] W. H. Organization, “Cardiovascular diseases,” last accessed 19 May 2024. [Online]. Available: https://www.who.int/health-topics/cardiovascular-diseases#tab=tab 1
[2] H. A. Skopicki and G. Gubernikoff, “Physical examination in the cardiac intensive care unit,” Cardiac Intensive Care (Second Edition), pp. 36–50, 2010.
[3] P. Carvalho, P. Gil, J. Henriques, L. Eug´enio, and M. Antunes, “Low complexity algorithm for heart sound segmentation using the variance fractal dimension,” IEEE International Workshop on Intelligent Signal Processing, 2005, pp. 194–199, 2005.
[4] S. Aziz, M. U. Khan, M. Alhaisoni, T. Akram, and M. Altaf, “Phonocardiogram signal processing for automatic diagnosis of congenital heart disorders through fusion of temporal and cepstral features,” Sensors, vol. 13, p. 3790, 2020.
[5] P. Varady, L. Wildt, Z. Benyo, and A. Hein, “An advanced method in fetal phonocardiography,” Computer Methods and Programs in Biomedicine, vol. 71, pp. 283–296, 2003.
References

## Slide 22
Thank you!

## Slide 23
Fabrication of the water-silicone acoustic transducer:
Appendix

## Slide 24
Discrete Wavelet Transform:
Appendix

## Slide 25
  Primary signal: Reference signal:
Noise-contaminated heart sound
Clean heart sound
Noise
Adaptive Noise Cancellation (ANC)
Dual Spectral Subtraction (DSS)
Correlated to
Appendix

## Slide 26
1. Extract the envelope of the denoised heart sound
2. Find the peaks S1 and S2
3. Find the boundaries of the peaks
Peaks
Start and end points of peaks
Appendix

## Slide 27
Appendix
Equations of the performance measurement metrics for denoising:
MSPE

## Slide 28
Appendix
Waveform of denoised heart sound with baseline wandering:

## Slide 29
Appendix
Summary of tuned parameters:

## Slide 30
Appendix
Parameter tuning for DWT:

## Slide 31
Appendix
Parameter tuning for ANC:

## Slide 32
Appendix
Emphasises the medium-amplitude signal and attenuates the low-amplitude signal.

## Slide 33
Appendix
Case where a noisy peak located closer to S1 than S2 is identified as an FHS. Ta < Tb < Tc results in the sequence [‘systole’, ‘systole’]

## Slide 34
Appendix
Decision rule for identification of missed FHS. Low-amplitude S2 is missed during the first-pass.

## Slide 35
Appendix
Decision rule for false peak removal. Noisy peak is identified as a peak in the first-pass.

## Slide 36
Appendix
Equations of the performance measurement metrics for segmentation:

## Slide 37
Primary signal:
Reference signal:
Estimated primary noise:
Methods
Error signal:
Primary noise:
Heart sound:
Update
Inputs: and
Filter estimates
Output:
Update the filter to approach
Methods: Adaptive Noise Cancellation
ANC in time domain

## Slide 38
Primary signal:
Reference signal:
Methods
Error signal:
Primary noise:
Heart sound:
Update
Estimated primary noise:
Inputs: and
Filter estimates
Output:
Update the filter to approach
Methods: Adaptive Noise Cancellation
ANC in time domain

## Slide 39
Primary signal:
Reference signal:
Estimated primary noise:
Denoised signal:
Methods
Primary noise:
Heart sound:
Update
Methods: Adaptive Noise Cancellation
ANC in time domain

## Slide 40
Example of DSS applied on noisy heart sounds:
Results
Objective 3
Results: Denoising (post tuning)